## Slide 1: The Vitek System
Lab 8

## Slide 2
The VITEK is an automated microbiology system utilizing growth –based technology , used for identification of m.o and antibiotic sensitivity test . The system is available in three formats (vitek 2 compact ,vitek 2, and vitek 2xL) that differ in increasing levels of capacity and automation . All three systems accommodate the same colorimetric reagent cards that are incubated and interpreted automatically
There are three types :-
VITEK 2
VITEK2compact
VITEK 2xL
Notes: I

## Slide 3: Reagent cards
The reagent cards have 64 wells that can each contain an individual test substrate . Substrates measure various metabolic activities such as acidification ,alkalinization ,enzyme hydrolysis ,and growth in the presence of inhibitory substances .An optically clear film present on both sides of the card allows for the appropriate level of oxygen transmission while maintaining a sealed vessel that prevents contact with the organism-substrate admixtures .each card has a pre inserted transfer tube used for inoculation .Cards have bar codes that contain information on product type ,lot number ,expiration date ,and a unique identifier that can be linked to the sample either before or after loading the card onto the system .

## Slide 4: There are four reagent cards :-
1- GN –Gram- negative fermenting and non fermenting bacilli
2- GP- Gram-positive cocci and non spore forming bacilli
3- YST –yeast and yeast –like organisms
4- BCL –Gram –positive spore –forming bacilli
Include acceptable culture media ,culture age ,incubation conditions ,and inoculum turbidity .
Culture Requirements

## Slide 5: Suspension preparation
A sterile swab or applicator stick is used to transfer a sufficient number of colonies of a pure culture and to suspend the microorganism in 3.0 ml of sterile saline (aqueous 0.4 % to 0.50% NaCl ,PH 4.5 to 7.0 ) in a 12 x 75 mm clear plastic (polystyrene) test tube . The turbidity is adjusted accordingly (see table 1) and measured using a turbidity meter called the DensiChek TM.
McFarland Turbidity Range | Product
0.50- 0.63 | GN
0.50-0.63 | GP
1.80-2.20 | YST
1.80-2.20 | BCL

## Slide 6: Identification cards are inoculated with microorganism suspensions using an integrated vacuum apparatus
1- A test tube containing the microorganism suspension is placed into a special rack (cassette) .
2- the identification card is placed in the neighboring slot while inserting the transfer tube the corresponding suspension tube .
3- the filled cassette is placed into a vacuum chamber station .
4- After the vacuum is applied and air is re-introduced into the station the organism suspension is forced through the transfer tube into micro- channels that fill all the test wells .